BIOTAGE ISOLERA – User Guide

- Book a slot and prepare the sample to load (samplet/dry loading if necessary)

- Switch on the instrument (the power switch is located underneath the touch screen)

- Ensure that all liquid lines are connected correctly and securely. Check the solvent and waste reservoirs’ volumes.

- Assign solvents to the solvent inlets (S1 and S2 are used for direct phase, S3 e S4 are used for inverse phase). Select:

>Chemistry_Setup_Solvent

- Without column, wash the lines with the solvents that are going to be used:

>Prime

- Without Cartridge;

- Volume 70 mL;

- Flow rate 40 mL/min;

- S1/S2 50:50

N.B. The PRIME function sends everything automatically to waste.

- Choose a method and a column:

>Method_Wizard

Follow the instructions on screen and proceed by clicking on NEXT.

Choose: user_sample name_solvent_TLC plate measurement_eluent_Rf product_Rf spot below_Rf spot above_column (choose a column between the ones proposed on screen depending on the sample weight and MOUNT IT)

- >RUN (column equilibration starts)

- When equilibration is completed CHARGE the sample:

>load sample

Once finished select:

>close

- To start the column:

>Gradient

- When the run is finished: wash the column with 100% of the most polar solvent - follow point 4 but ‘With Cartridge’. To dry the column select AIR FLUSH.

- Take out the column and wash the lines with methanol – follow point 4, without cartridge, putting both lines in methanol and selecting S1/S2 50:50

- (Optional) Save on a USB drive the purification report:

>Results_Archive Report.

The results are saved as .pdf files.

- Switch off the instrument from the main menu by selecting “Shut Down” e only after press the ON/OFF power switch under the touch screen.

PLEASE REMEMBER

- TO ALWAYS FILL THE REGISTER (date, hour, solvents used, operator)

- At the end of the purification, LEAVE THE SOLVENTS AS FOLLOWS:

S1/S2: DIRECT PHASE

S3: ACETONITRILE

S4: ACQUA

- UV LAMP:

monitor range 200 – 400 nm

UV1 = 254 nm

UV2 = 210 nm

- ALWAY CHECK THAT “General_Collection all_ON” IS SELECTED to collect all fractions. It’s possible to leave it OFF only if the purification has already been tested and the elution of the solvent is well known.

- Biotage SHOULDN’T BE USED for routine purifications. Columns are expensive and theoretically single use, so please use it only for difficult or inverse phase purifications.