HOW TO KNOW ABOUT COLUMNS

Each column has a specific range of stability about pressure, flow rate, pH and temperature. Before to mount a column please always check this information on dedicated table.

For the optimal method and good maintenance of columns, also check the maximum amount of sample, injection volume and minimum wash volume on the same table. You should remind to filter all your sample with appropriate HPLC filters (0.45 µm), especially when the column hasn’t got a guard column! Before the journal shutdown of system or before removing the column, perform the final washing procedure:

- W/ACN 90/10, 1 mL/min for 15 min;

- W/ACN 10/90, 1 mL/min for 15 min.

Store the columns in a mixture of organic solvent and water (e.g. W/ACN 25/75) after final washing procedure and put the caps at the end. Never store the columns with some buffer inside (it request different washing procedure), because this will short the lifetime.

Equilibration (for the beginning)

Cartridges need to be flushed properly before the use. Starting from a flow equals to 0, gradually increase the flow rate in small manual steps, or using the settable gradient directly into the software (Pump > Auxiliary… > set at 1 mL/min). This method must always be used when the flow in column starts from zero, this avoids a strong impact on the stationary phase preventing its breakage.

If the column has dried out during the storage, follow this special method (during the flux please check frequently the pressure profile):

- Activate the packing by flushing W/ACN 5/95 at 0.1 mL/min for 1 or 2 hours;

- Increase flow at 0.3 mL/min for 30 min;

- Increase flow at 0.5 mL/min for 30-20 min;

- Increase flow at 1 mL/min for 10 min.

You can change timing by evaluating the UV line at 210 nm and 254 nm, when the profile become dish you can pass at next step. If all went well, now the column is packed and you can proceed with the equilibration.

Adapt the method in case of different maximum flow rate, long length or if the column is a direct phase.

Reversed phase columns (RP-18, RP-8) are shipped in W/ACN (50/50 or 25/75) and normal phase columns (Si, NH2, CN, Diol) are shipped with n-heptane/dioxane 99/1, or differently if specified.

Procedure for equilibration:

- Set flow at 1 mL/min with W/ACN 10/90 for 10 min;

- Set your initial eluent of the method at 1 mL/min for 10/15 min;

adjust the time depending on the column length

- Leave running the eluent at flow of your method, the injection should start with the pump on.

Now the column is ready for the injection.

Washing program (for the end of method)

At the end of each method, program the wash of column as follow (the time depends on the column length):

- Bring the eluent to W/ACN 5/95 in gradient;

- Isocratic at this composition;

- Gradient to W/ACN 90/10;

- Isocratic W/ACN 90/10.

If the UV and pressure lines are stable, the column is ready for a new injection.

GUIDE TO GOOD USE HPLC

Fill Register USERS.

Open purge pump with a little rotation (you should feel a friction decrease), turn on the system using all physical buttons, and open the Agilent ChemStation software.

Choose your solvents and Check:

- the eluent volume into the bottle;

- dip-pipes are immersed and close to the bottom;

- no air in all lines;

- the method opened.

Select DEF method. Don’t modify any method other than yours! DEF is preset and cannot be overwritten, so

you can operate in security.

Set, or check, the real volume of each eluent into the software (click on Bottle image > Solvent bottle filling

> fill the boxes with volume and the lower level admitted before stopping flow) Click on Pump > Auxiliary > Set properly compressibility

If in the line is present some air, connect the capillary instead of a column, close purge pump, click on Pump > Set pump > 1 mL/min; set 100% of line that contains air and keep this set for 15 min or more (you should check the waste line, if you see bubbles continue with this conditions). When pressure become stable the system is ready to proceed.

If in the lines is present W/i-PrOH eluent, keep install the capillary, leave purge pump open, and run the correct eluent clicking on Pump > Set pump > 5 mL/min; set properly the percentage of solvent and condition the lines keeping this set for 15 min. Do this for the lines you will need. Stop the flow, close the purge pump, click again on Pump > Set pump > 1 mL/min and keep conditions for other 5 min. Then you can proceed.

Turn on the pump and click on Pump > Set pump > 5 mL/min; Name of solvent; set 50% of B and condition the line keeping this set for 10-15 min (skip this part if you already performed the previous procedure in blue box).

Stop the flow and click on Pump > Auxiliary > set 1.00 mL/min as gradient flow.

If there isn’t air in the system, put your column, close purge pump, turn on DAD (> Control > ON only lamp you will need)

Equilibrate the column: click on Pump > Set up pump > fill the box with “Procedure for equilibration”

Set you method in pump and DAD (or open it, if so check again Pump > Auxiliary > 1 mL/min), remember to

add “Washing program (for the end of method)”.

Click on Injector > Set up Inj > select Inj with needle wash and fill the box with Vial 91 (if the Vial isn’t already

present, put in one in this position containing MeOH).

Remember to set Standby mode of system at the end of a Sequence (Sequence > Sequence Parameter > tick on Post-Sequence Command/Macro > Standby and set Not Ready Timeout at 2 min) or if no one will use the instrument after then you (for manual control: Pump > Control.. > Standby)

Procedure to shutdown the system:

- At the end of each day proceed with the “final washing procedure” (with the column connect: W/ACN 90/10, 1 mL/min for 15 min, then W/ACN 10/90, 1 mL/min for 15 min);

- At the end of the week, or before a long stop period, wash the instrument

  - perform the “final washing procedure”,

  - open purge pump,

  - replace the column with the capillary,

  - put all dip-pipes into a solution of W/i-PrOH 1:1,

  - set the pump with A1 active, B1 = 50%, 5 mL/min for 10 min,

  - set the pump with A2 active, B2 = 50%, 5 mL/min for 10 min,

  - stop flow, close purge pump and set flow at 1 mL/min for 5-10 min.

Proceed with turn off the modules (pump and DAD) on software, close software, and turn off the buttons.

Remember that if the lines of HPLC contain W, ACN, or other organic solvents, turn on the instrument and eluent running for 15 min in each lines every day to avoid the degradation lines. Finish this procedure with “final washing procedure”.

REGISTER USERS

| DATE | NAME | COLUMN TYPE | SOLVENTS USED | SIGN |
| --- | --- | --- | --- | --- |
|  |  |  | A1 | A2 | B1 | B2 |  |

REGISTER WEEKLY WASH INSTUMENT W/i-PrOH

| DATE | NAME | SIGN | DATE | NAME | SIGN |
| --- | --- | --- | --- | --- | --- |